Table 1. Method-comparison metrics for the four static scenarios (n = 150 paired measurements each).

| Scenario | Mean bias (mmHg) | 95% LoA (mmHg) | PE (%) | BA reg. slope (95% CI) | Deming slope (95% CI) | P-B slope (95% CI) | CCC | C_b | v |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| S1 offset only | +11.9 | 5.6 to 18.2 | 4.6 | 0.010 (-0.007, 0.028) | 1.010 (0.993, 1.028) | 1.010 (0.990, 1.029) | 0.921 | 0.926 | 1.010 |
| S2 zeroed (ideal) | +0.1 | -5.4 to 5.5 | 4.2 | -0.010 (-0.025, 0.006) | 0.990 (0.975, 1.006) | 0.985 (0.971, 1.001) | 0.995 | 1.000 | 0.990 |
| S3 gain error | +12.7 | 4.7 to 20.7 | 5.9 | 0.094 (0.080, 0.109) | 1.099 (1.082, 1.117) | 1.100 (1.081, 1.117) | 0.915 | 0.919 | 1.099 |
| S4 gain masked | +0.3 | -8.4 to 8.9 | 6.7 | 0.103 (0.087, 0.119) | 1.109 (1.091, 1.126) | 1.111 (1.092, 1.130) | 0.990 | 0.995 | 1.108 |

BA = Bland-Altman; CI = confidence interval; LoA = limits of agreement; PE = percentage error; P-B = Passing-Bablok; CCC = concordance correlation coefficient; C_b = bias-correction factor; v = scale shift.

Table 2. Dynamic-response metrics for pulse pressure (PP) and systolic pressure (SBP).

| System (f_n, zeta) | PP ratio | PP mean bias (mmHg) | CCC (PP) | v (PP) | SBP mean bias (mmHg) |
| --- | --- | --- | --- | --- | --- |
| Optimal (25 Hz, 0.65) | 1.01 | +0.3 | 0.973 | 1.030 | +0.2 |
| Underdamped (10 Hz, 0.15) | 1.24 | +10.7 | 0.588 | 1.277 | +7.6 |
| Overdamped (8 Hz, 0.80) | 0.92 | -3.4 | 0.891 | 0.989 | -3.4 |

PP ratio = mean(measured PP / true PP); f_n = natural frequency; zeta = damping coefficient.